## Sheet1
| CALIFORNIA SOCIETY of ENVIRONMENTAL ANALYSTS Laboratory Internal Audit |  | SM 9221 B, C, E, F
SOP: |  |  | Audit date:
Analysts audited: |  |  | Auditor:
Method Reference | Requirement | Conforms | Notes and Observations
Medium, Reagents, and Standards
9050.A.1 | Dehydrated media is stored in a cool dry location away from light at less than 30oC and discarded by the manufacturer’s expiration date.
9050.A.1 | Media prepared commercially is discarded by manufacturer’s expiration date.
Lauryl tryptose broth
Used for 9221.B.3 Presumptive Phase
9221.B.3.a | Alternatively, omit inverted vial and add 0.01g/L bromocresol purple to presumptive medium to determine acid production, the indicator of a positive result in this part of the coliform test.
9221.B.3.a | pH should be 6.8 + 0.2 after sterilization
9221.B.3.a | Make lauryl tryptose broth of such strength that adding 100mL, 20-mL, or 10-mL portions of sample to medium will not reduce ingredient concentrations below standard medium.
9020.B.5.2.j | Stored at 15°-25°C in the dark for no more than 3 months
9020.B.5.2.j | Discard unused medium after 3 months or sooner if there is evidence of moisture loss, medium contamination, or medium deterioration (darkening of the medium)
Brilliant green lactose bile broth
Used for 9221.B.2 Conﬁrmed Phase
9221.B.4.a | pH should be 7.2 + 0.2 after sterilization
9020.B.5.2.j | Stored at 15°-25°C in the dark for no more than 3 months
9020.B.5.2.j | Discard unused medium after 3 months or sooner if there is evidence of moisture loss, medium contamination, or medium deterioration (darkening of the medium)
EC medium
Used for 9221 E. Fecal Coliform Procedure
9221.E.1.a | pH should be 6.9 + 0.2 after sterilization
9221.E.1.a | Sterilize by autoclaving at 121°C for 12 to 15 min
9050 A.1. | Do not expose tubes to direct sunlight; store at <30°C.
9020.B.5.2.j | Discard unused medium after 3 months or sooner if there is evidence of moisture loss, medium contamination, or medium deterioration (darkening of the medium)
A-1 broth
Used for 9221 E. Fecal Coliform Procedure Direct Isolation
9221.E.2.a | Heat to dissolve solid ingredients and add polyethylene glycol p-isooctylphenyl ether
9221.E.2.a | Adjust to pH 6.9 + 0.1
9221.E.2.a | Sterilize by autoclaving at 121°C for 10 min.
9221.E.2.a | Store in dark at room temp for not longer than 7 days.
9221.E.2.a | Make A-1 broth of such strength that adding 10-mL sample portions to medium will not reduce ingredient concentrations below those of the standard medium. For 10-mL samples prepare double-strength medium.
EC-MUG medium
Used for 9221 F. Escherichia coli Procedure
9221.F.1.a | Dispense in tubes that do not fluoresce under long-wavelength (365-366nm) ultraviolet (UV) light
9221.F.1.a | Sterilize for 15 min at 121°C.
9221.F.1.a | pH should be 6.9 ± 0.2 after sterilization
GAD reagent
Used for 9221 G. Other Escherichia coli Procedures
9221.G.1.a | pH should be 3.4 + 0.2
9221.G.1.a | Stable for 2 months when stored at 5° C.
9221.G.1.a | It can be ﬁlter-sterilized (0.2um filter) and treated as a sterile solution.
Tryptone water
Used for 9221 G. Other Escherichia coli Procedures
9221.G.2.a | Adjust pH to 7.5.
9221.G.2.a | Sterilize for 10 min at 121°C.
Kovacs’ reagent
Used for 9221 G. Other Escherichia coli Procedures
9221.G.2.a.2 | Dissolve aldehyde in alcohol
9221.G.2.a.2 | Store in the dark at 4°C.
9221.G.2.a.2 | This reagent should be pale yellow to light brown in color. Use of low-quality amyl alcohol may produce a dark-colored reagent; do not use such a reagent.
Dilution Water
9050.C.1 | Stock buffer solution or peptone water is prepared as specified in Standard Methods 9050C
9050.C.1 | Stock buffers are autoclaved or filter-sterilized
9050.C.1 | Stock solutions are discarded if turbidity develops.
9020.B.9.d | Each batch (or lot, if commercially prepared) of dilution/rinse water is checked for sterility by adding 50 mL of water to 50 mL double strength non-selective broth, incubating at 35o ± 0.5oC, and checking for growth after 24 hours and 48 hours
9050.C.1 | Is the batch/lot discarded if growth is detected
Sample Collection and Preservation
9060.A.5 | The laboratory measures and records the sample temperature upon arrival
9060.A.5 | The laboratory rejects samples when the paperwork, preservatives, containers and holding times are incorrect or cannot be verified
9060.B | The laboratory complies with the preservation and holding times for the analytes of interest.
9060.A.2 | Samples dechlorinated using Na2S2O3
9060.A.3 | Was 100mL collected in sample container with at least 2.5cm (1 inch) of space above sample to facilitate mixing?
9060.A.3 | If the container does not have enough headspace to allow proper mixing, is the entire sample transferred to a larger, sterile container, mixed properly, then measured to ensure 100mL of sample is analyzed?
9060.B | Maximum elapsed time between collection and analysis of samples is 8-30 h depending upon rule.
Quality Control
The laboratory performed an Initial Demonstration of Capability for each analyst and organism.
The laboratory performs Ongoing Demonstration of Capability for each analyst and organism.
The laboratory shall perform a sterility check for each lot of prepared, ready-to-use, media and on each batch of media prepared in the laboratory.
The laboratory shall conduct method blanks per the analytical method.
All growth and recovery media shall be checked to assure that the target organism(s) respond in an acceptable and predictable manner once per lot or batch.
9020.B.9.b | In order to ensure identity and traceability, reference cultures used for positive and negative controls shall be obtained from a recognized national collection, organization, or manufacturer recognized by the accreditation body.
9020.B.9.b | Working stocks shall not be sequentially cultured more than five (5) times and shall not be sub-cultured to replace reference stocks.
9020.B.9.b | Each pre-prepared, ready-to-use lot of medium and each batch of medium prepared in the laboratory shall be tested with at least one (1) positive and one (1) negative culture controls and that produce typical results based on the method. This shall be done prior to first use of the medium.
The laboratory participates in PT studies at least twice per year if they are TNI accredited, otherwise once per year.
Procedure
9221 B. Standard Total Coliform Fermentation Technique
9221.B.3.b | Mark each dilution set with sample number, dilution, date, and any other necessary information.
9221B.3.b.1 | Thoroughly mix all samples or dilutions by rapidly making about 25 complete up-and-down (or back-and-forth) movements or shake vigorously for 5 s.
9221B.3.b.1 | For potable water use ﬁve 20-mL portions, ten 10-mL portions, or a single bottle of 100 mL portion; for nonpotable water use ﬁve tubes per dilution of 10, 1, 0.1 mL.
9221B.3.b.1 | Inoculate each tube in a set of ﬁve with replicate sample volumes. Mix test portions in the medium by gentle agitation.
9221B.3.b.2 | Incubate inoculated tubes or bottles at 35 + 0.5°C. After 24 + 2 h swirl each tube or bottle gently and examine it for growth, gas, and acidic reaction (shades of yellow color) and, if no gas or acidic reaction is evident, reincubate and reexamine at the end of 48 + 3 h.
9221B.3.c | Submit all presumptive tubes or bottles showing growth, any amount of gas, or acidic reaction to the conﬁrmed phase.
9221B.4.b | Inoculate confirmation tubes by transfering one or more loopfuls of culture to a fermentation tube containing brilliant green lactose bile broth or insert a sterile wooden applicator to transfer culture.
9221B.4.b | Incubate the inoculated brilliant green lactose bile broth tube at 35 + 0.5°C.
9221B.4.b | Formation of gas in any amount in the inverted vial of the brilliant green lactose bile broth fermentation tube at any time within 48 + 3 h constitutes a positive conﬁrmed phase.
9221B.4.c | Alternative procedure may only be used for polluted water or wastewater. If all presumptive tubes are positive in two or more consecutive dilutions within 24 h, submit to the conﬁrmed phase only the tubes of the highest dilution (smallest sample inoculum) in which all tubes are positive and any positive tubes in still higher dilutions. Submit to the conﬁrmed phase all tubes in which gas or acidic growth is produced only in 24 to 48 h.
9221 E. Fecal Coliform Procedure
9221.E.1.b.1 | Submit presumptive fermentation tubes or bottles showing any amount of gas, growth, or acidity within 48 h of incubation to the fecal coliform EC medium test.
9221.E.1.b.1 | Gently shake or rotate presumptive fermentation tubes or bottles showing gas, growth, or acidity. Using a sterile loop or sterile wooden applicator stick, transfer growth from each presumptive fermentation tube or bottle to EC broth.
9221.E.1.b.2 | Incubate inoculated EC broth tubes in a water bath at 44.5 + 0.2°C for 24 + 2 h.
9221.E.1.b.2 | Place all EC tubes in water bath within 30 min after inoculation. Maintain a sufﬁcient water depth in water bath incubator to immerse tubes to upper level of the medium.
9221.F.2 | A-1 broth Direct Isolation of fecal coliforms:
9221E.2.b. | Thoroughly mix all samples or dilutions by rapidly making about 25 complete up-and-down (or back-and-forth) movements or shake vigorously for 5 s.
9221E.2.b. | Inoculate each tube in a set of ﬁve with replicate sample volumes. Mix test portions in the medium by gentle agitation.
9221E.2.b. | Incubate for 3 h at 35 ± 0.5°C then transfer tubes to a water bath at 44.5 ± 0.2°C and incubate for an additional 21± 2 h.
9221 F. Escherichia coli Procedure
9221.F.1 | Submit all presumptive fermentation tubes or bottles from 9221B or 9221D showing growth, gas, or acidity within 48 + 3 h of incubation to the E. coli test.
9221.F.1.b.1 | Gently shake or rotate presumptive fermentation tubes or bottles showing growth, gas, or acidity. Using a sterile 3- or 3.5-mm-diam metal loop or sterile wooden applicator stick, transfer growth from presumptive fermentation tube or bottle to EC-MUG broth.
9221.F.1.b.2 | Incubate inoculated EC-MUG tubes in a water bath maintained at 44.5 + 0.2°C for 24 + 2 h. Place all EC-MUG tubes in water bath within 30 min after inoculation.
9221.F.1.b.2 | Maintain sufficient water depth in the water-bath incubator to immerse tubes to the upper level of medium.
9221.G | For E. coli determination using GAD Medium: Submit all presumptive fermentation tubes or bottles from 9221B or 9221D showing growth, gas, or acidity within 48 + 3 h of incubation to the E. coli test using GAD medium.
9221.G.1.b.1 | Gently shake or rotate presumptive tubes or bottles showing growth, gas, or acidity. Using a graduated pipet, transfer 5 mL broth from the fermentation tube or bottle to 15-mL centrifuge tube.
9221.G.1.b.2 | Concentrate the bacterial cells from the broth by centrifugation for 10 min. Discard supernatant and resuspend cells in 5 mL phosphate buffer. Reconcentrate cells by centrifugation for 10 min. Discard supernatant and add 1.0 mL GAD reagent. Vigorously swirl the tube to resuspend cells in GAD reagent.
9221.G.1.b.3 | Incubate tubes at 35°C and observe after 1 h. Tubes may be incubated for a maximum of 4 h.
9221.G.2 | For E. coli determination using tryptone water and Kovacs’ reagent submit all presumptive fermentation tubes or bottles from 9221B showing growth, gas, or acidity within 48 +3h of incubation to this test.
9221.G.2.b | Using a sterile metal loop or sterile wooden applicator stick, transfer growth from presumptive fermentation tube or bottle to a tube containing 5 mL tryptone water.
9221.G.2.b | Incubate inoculated tryptone water tubes in a water bath or incubator maintained at 44.5°C for 24 ± 2 h.
9221.G.2.b | After incubation, add 0.2 to 0.3 mL Kovacs’ reagent to each tube of tryptone water.
Data Analysis and Calculations
9221 B. Standard Total Coliform
9221B.3.b.2 | Record presence or absence of growth, gas, and acid production. If the inner vial is omitted, growth with acidity signiﬁes a positive presumptive reaction.
9221.B.4 | Simultaneous inoculation into brilliant green lactose bile broth for total coliforms and EC broth for fecal coliforms or EC-MUG broth for Escherichia coli. Positive brilliant green lactose bile broth cultures (35°C) and positive EC and EC-MUG broths at elevated temperature (44.5°C) results as a positive completed test response.
9221.B.4 | Positive brilliant green lactose bile broth cultures with negative EC or EC-MUG broth cultures indicate the presence of nonfecal coliforms.
9221.C | Enumeration of results reported using MPN table in SM 9221C Estimation of Bacterial Density.
9221 E. Fecal Coliform
9221.E.1.c. | Gas production with growth in an EC broth culture within 24 + 2 h or less is considered a positive fecal coliform reaction. Failure to produce gas (with little or no growth) constitutes a negative reaction.
9221.E.1.c. | If multiple tubes are used, calculate MPN from the number of positive EC broth tubes as described in Section 9221C.
9221.E.1.c. | When using only one tube for subculturing from a single presumptive bottle, report as presence or absence of fecal coliforms.
9221.E.2.c | Gas production in any A-1 broth culture within 24 h or less is a positive reaction indicating the presence of fecal coliforms.
9221.E.2.c and 9221.C | Enumeration of results reported using MPN table in SM 9221C Estimation of Bacterial Density.
9221 F. Escherichia coli
9221.F.1.c | When using EC-MUG for the examination of E. coli. Examine all tubes exhibiting growth for ﬂuorescence using a long-wavelength 365-366 nm UV lamp in darkened room. The presence of bright blue ﬂuorescence is considered a positive response for E. coli.
9221.G.1.c | When using GAD reagent for the conﬁrmation of E. coli. Examine all tubes for a distinct color change from yellow to blue. The presence of a blue color is considered a positive response for E. coli.
9221.G.2.c | When using tryptone water and Kovacs’ reagent for conﬁrmation of E. coli. Examine all tubes for appearance of a deep red color in the upper layer. The presence of a red color is considered a positive response for E. coli.
9221.G.3 and 9221.C | Enumeration of results reported using MPN table in SM 9221C Estimation of Bacterial Density.
Drinking Water Requirements EPA Certification of Laboratories Analyzing Drinking Water, Fifth edition
6.3.3 | Sample transported on ice and maintained below 10°C unless the time since sample collection has been less than two hours.
5.3.1.2.2. | The media must be protected from light.
5.1.6.2 | For media prepared in the laboratory is the following information recorded? - Date of preparation - Type of medium - Lot number - Sterilization time and temperature - Final pH (after sterilization) - Technician’s initials
5.1.6.3 | For media prepared commercially is the following recorded for each lot? - Date received - Type of medium - Lot number - pH verification
5.1.6.4 | Is each new lot of dehydrated or prepared commercial medium and each batch of laboratory-prepared medium checked before use for sterility and with positive and negative culture controls?
5.2.2.1 | For drinking water samples, is the total sample volume of 100 mL used for each test?
5.2.2.2 | For source water samples, are at least 3 series of five tubes each with appropriate sample dilutions used?
5.2.2.7.2 | For source water samples, are all samples that produce a turbid culture (i.e., heavy growth) in the absence of gas/acid production in LTB or LB invalidated?
5.2.2.6 | For drinking water samples, is each total coliform-positive sample tested for the presence of either fecal coliforms or E. coli?
5.2.2.7.1 | For drinking water samples, are all samples that produce a turbid culture (i.e., heavy growth) in the absence of gas/acid production, in LTB or LB, invalidated?
5.2.2.7.2 | For source water samples, are all samples that produce a turbid culture (i.e., heavy growth) in the absence of gas/acid production in LTB or LB invalidated?
5.2.3.1.1 | When six-times formulation strength medium is used, is it filtersterilized rather than autoclaved?
5.2.3.4 | Are yellow cultures confirmed in BGLBB and a fecal coliform/E. coli test conducted?
5.2.3.5 | Are all samples which produce a non-yellow turbid culture in P-A medium invalidated?
5.2.4 | Fecal Coliform Test uses EC Medium for fecal coliforms in drinking or source water, or A-1 Medium for fecal coliforms in source water only.
5.2.4.1.1 | Is EC medium used to test a total coliform-positive culture for fecal coliforms under the Total Coliform Rule?
5.2.4.1.4 | Are the inverted vials examined to ensure that they are free of air bubbles and at least one-half to two-thirds covered after the sample is added?
5.2.4.2.1 | If A-1 Medium is used, is it used to enumerate only fecal coliforms in source water, in accordance with SWTR, and not for drinking water samples?
5.2.4.2.1 | Are three sample volumes of source water used in a five- or ten tube/sample volume format?
5.2.4.2.3 | Are inverted tubes examined to ensure that they are free of air bubbles?
5.2.4.2.5 | A-1 broth may be held for up to 7 days in a tightly closed screw-cap tube at 4°C.
HOLDING TIME and SAMPLE PRESERVATION REQUIREMENTS
Analyte | Holding Time
Total Coliform | 8-30 hours Depending on Source
Fecal Coliform | 8-30 hours Depending on Source